## Sheet1
certificate_no | N21080387/D6.02c/C-06
device_name | DC REFERENCE STANDARD
end_date | 21.09 2021
next_date | 21.03.2000
calibrated_for | LF, HF Impedance and DC Metrology CSIR-National Physical Laboratory. Dr. K.S. Krishnan Road, New Delhi- 110012 Customer Reference: Note dated 15.07.2021
description | DC Reference Standard Model No.: 732B S..No.: 9130702 Make: Fluke Corporation. USA
env_conditions | Temperature: (25 ± 1)℃ Humidity (RH): (50 ± 10)%
stds_used | Programmable Josephson Voltage Standard. NA (Quantum based Primary Voltage Standard )
tracability | Programmable Josephson Voltage Standard. (Primary standard)
procedure | The DC Reference standard has been calibrated by comparison method with Programmable Josephson Voltage Standard as per calibration procedure no. 6.02c/Doe.3/CP # 01.
calibrated_by | (SANDHYA M. PATEL)
checked_by | (SAOOD AHMAD)
incharge | (ANURAG GUPTA)
issued_by

## Sheet2
result_table | S.no. | Nominal Value | Measured value | Expanded Uncertainty
| 1 | 10.0 v | 10.000 038 35 v | = 91 n V
| 2 | 1.018V | 1.018 159 396 v | 86 n V
result_desc | The reported expanded uncertainty is at coverage factor k = 2 which corresponds to a coverage probability of approximately 95% for a normal distribution.
calibration_date | 24th August to 27th August 2021.
remarks | (i) The DC Reference standard has been calibrated in AC power ON condition. (ii) The noise of the DC reference Standard is inclusive in the estimation of the uncertainty in measurement.